## Slide 1
W. Linwood Jones and Cathy May
Central Florida Remote Sensing Laboratory
University of Central Florida
Christopher S. Ruf
Space Physics Research Laboratory,
University of Michigan
Observations of C-Band Brightness Temperature from the Hurricane Imaging Radiometer (HIRAD) on-board the NASA WB-57 during GRIP in Hurricanes Earl And Karl (2010)
HIRAD Technology Investment Roadmap
Earl Flight: 1-2 Sept 2010
Timothy Miller, Mark James, J. Brent Roberts,
Sayak Biswas and Courtney Buckley
NASA Marshall Space Flight Center
Eric Uhlhorn
Atlantic Oceanographic and Meteorological
Laboratory (AOML)
Peter Black
SAIC, Inc at NRL/
Marine Meteorology Division
HIRAD first flight- 2010: NASA GRIP- Genesis and Rapid Intensification Project (WB-57)
Earl and Karl Storm Tracks: HIRAD WB-57 Flight Locations
Karl 16 Sept 2010
Next up: NASA Hurricane and Severe Storm Sentinel (HS3) mission 2012-14 (Global Hawk)
Key features of radar and HIRAD comparison (right):
  1) Eyewall dBZ and HIRAD Tb maxima are located in left-front quadrant. This, together with open eyewall in right-rear quadrant, suggest easterly environmental shear.
  2) Partial outer concentric eyewall is suggested equally well in radar and HIRAD images, as is outer principal rainband.
Operational use of SFMR for Ocean Surface Wind Speed and Rain Rate since 2002 & since 1980 NASA LaRC AFE/RADSCAT flights
(AF WC-130J, NOAA WP-3D)
Improved Hurricane Ocean Surface Vector Winds and Rain Rate
Technology Transfer: Operational Tropical Cyclone Reconnaissance
HIRAD flight
Hurricane Earl
Sept
Aug
HIRAD flights
Best day
Hurricane Karl
Sept
GRIP (2010) HIRAD
Swath Geometry
Similar to NASA HS3 Global Hawk
  configuration for 2012-2014
“Model” data are Tb’s computed from SFMR wind & rain fields
SFMR
SFMR
HIRAD 4- and
5-GHz observed
Tb values agree
with SFMR Tb
values
HIRAD obs from
WB-57 at 20-km;
SFMR obs from
WP-3D at 3 km
Model Tb@ 4GHz
HIRAD Tb@ 4GHz
HIRAD Tb@ 5GHz
Model Tb@ 5GHz
SFMR Wind +50 (m/s)
SFMR Rain +50 (mm/hr)
50
30
10
WS (kt)
Tb (K)
30
10
RR (mm/hr)
(kt)
(mm/hr)
Hurricane Imaging RADiometer (HIRAD)
HIRAD Passive Microwave C-Band Radiometer :
Version 1: H-pol at 4, 5, 6 & 6.6 GHz freq for wind speed and rain rate
Version 2: H- & V-pol continuum of C-band freq for Ocean Surface Vector Wind (OSVW) & rain rate
Performance Characteristics (Version 1):
Earth Incidence angle: 0- 60°,
Spatial Resolution: 2-5 km,
Swath: ~70 km for 20 km altitude
Observational Goals (Version 1):
WS: 10 to >85 m/s; RR: 2 to > 100 mm/hr
HIRAD advances:
Expands Stepped Frequency Microwave Radiometer (SFMR) coverage from nadir to wide swath
Eliminates SFMR max wind under-sampling with wide-swath sampling of entire eyewall (instantaneously)
Uses synthetic thinned array and RFI mitigation technology of Lightweight Rain Radiometer (NASA Instrument Incubator)
HIRAD
Version 1
HIRAD
Version 2
Version 2: A fully polarimetric continuum of frequencies will be measured from which 6 will be sampled that will result in wind direction retrievals as well as more accurate wind speeds (fields of OSVW)
This approach will eliminate wind and rain retrieval ambiguities in gale-force wind regimes.
Chart: Karl 16 Sept Center Position
| -93.343000000000004 | -93.458000000000013 | -93.617999999999995 | -93.99
P3 reported | 19.664 | 19.706 | 19.767 | 19.713
HIRAD subjective | 19.74 | 19.77 | 19.8 | 19.819 | 19.8
P3 radar subjective | 19.684 | 19.792 | 19.808
Good agreement between HIRAD
  and WP-3D LF radar centers
WP-3D wind centers consistently south
  of LF radar and HIRAD centers as seen
  in radar composite image to right
HIRAD 5 GHz excess Tb (K)
Storm track
Storm Track
60
-60
0
60
0
-60
HIRAD GRIP Calibration Challenges
Rain rate and wind speed retrievals to be based on three calibrated frequencies:
  5 GHz Tb (Microwave Brightness Temps) have been successfully produced
  Calibration of other 2 channels is in progress; completion within 3 months
HIRAD calibration issues and mitigation for HS3:
  Calibration uses internal reference blackbody targets and noise diodes
  Calibration algorithm depends on reference Tb: uncorrected instrument temperature dependence found during GRIP flights (+/-25o C)
  Temperature correction algorithm being developed for GRIP (requires additional instrument characterization testing)
  Thermal control subsystem has been upgraded for HS3 to minimize instrument temperature fluctuations
HIRAD Physical Principles
HIRAD measures ‘emissivity’ over a range of incidence angles and frequencies from ocean surface foam (a function of wind speed) and intervening rain.
At C-band microwave frequencies (4-6.6 GHz), emissivity from wind-driven foam is invariant with frequency while rainfall emissivity is a strong function of frequency.
These physical characteristics allow two geophysical variables (wind speed and rain rate) to be derived from emissivity measurements at 4-6 discrete C-band frequencies, an ‘over-determined’, ‘least-squares’ problem solvable with conventional mathematical techniques.
Surface wind speed and rain rate retrievals are derived from empirical correlation of HIRAD measured emissivity at operating incidence angles with co-located GPS dropsonde surface wind observations and accepted functional relations for rain attenuation vs frequency.
WP-3D LF radar reflectivity dBZ
WB-57 flight track
P-3 composite center time ~20:44 UTC
Radar; + HIRAD; + Wind Centers
Partial outer concentric eyewall
0
60
120
180
-60
-120
0
60
120
-60
-120
Outer principal rainband